## Sheet1
Phase A
We consider In=10A, Vn=230V, MC=8000imp/kwh, eval board voltage divider
| fL[Hz] |  | 50
| In[A] |  | 10
| Vn[V] |  | 230
| MC[imp/kwh] |  | 8000
| Full scale voltage [V] | full scale voltage, VFS=0.5/sqrt(2)*901 for evaluation board | 318.55
| Full scale current [A] |  | 120
| pf |  | 1
| ϕ=acos(pf) |  | 0
| CF active power expected at In, Vn, pf, CF=MC*V*I*PF/3.6E6 |  | 5.1111
| CF reactive power expected at In, Vn , pf, CF=MC*V*I*sin(ϕ)/3.6E6 |  | 0
| Vn as percetage of full scale |  | 72.2022
| In as percetage of full scale |  | 8.3333
| CF at full scale |  | 68800
| CFDEN=CF_FS*V%*I%*pf/CF_exp |  | 810
| VLEVEL=VFS/Vn*4E6 |  | 5540000
| PMAX |  | 26991271
| LINECYC |  | 100
| PERIOD |  | 5122
| fs[Hz] |  | 1024000
| WTHR |  | 3
PHASE COMPENSATION
| pf=0.5 |  | 0.5
| ϕ=acos(pf) |  | 1.0472
| Expected WATTHR |  | 2066
| Expected VARHR |  | 3578
| AWATTHR |  | 2042
| AVARHR |  | 3402
| error=-atan[(WATTHR*sin(ϕ)-VARHR*cos(ϕ))/(VARHR*sin(ϕ)+WATTHR*cos(ϕ)) |  | -0.97366
| phase resolution=360*fL/1.024E6 |  | 0.017578
| Phase Compensation=abs(Error/Phase resolution) |  | 55
| APHCAL |  | 55
| AWATTHR after APHCAL compensation |  | 1984
| AVARHR after PHCAL compensation |  | 3436
| error between AWATTHR just obtained and the expected one[%] |  | -3.969
| error between AVARHR just obtained and the expected one[%] |  | -3.9687
| error=-atan[(AWATTHR*sin(ϕ)-AVARHR*cos(ϕ))/(AVARHR*sin(ϕ)+AWATTHR*cos(ϕ)) |  | -0.0028073
GAIN CALIBRATION
| APGAIN=WATTHRexpected/WATTHRactual-1 |  | 346707
| AWATTHR after APHCAL and APGAIN compensation |  | 2066
| AVARHR after APHCAL and APGAIN compensation |  | 3578
| error between AWATTHR just obtained and the expected one[%] |  | 0
| error between AVARHR just obtained and the expected one[%] |  | 0